3.8 DC Transistorized  voltmeter

3.9. Introduction

The DC transistorized voltmeter (TVM) is composed of a TVM attenuator and a null detector. The primary part of the TVM is the null detector, in wich the DC signal is modulated by a photochopper, amplified by a five-stage amplifier, rectified by a phase-sensitive demodulator, and filtered to produce a DC output. Transformer coupling is used to isolate the DC meter/recorder output from the floating null detector circuitry. Negative feedback makes the null detector insensitive to gain variations of individual transistors. The panel meter has tautband suspension, which eliminates the stickiness associated with pivot-jewel meters.

3.10. Null-detector

The null detector is a current feedback amplifier that drives a zero center meter. The feedback voltage is the voltage drop across R208 caused by a portion of the output current. At the input to the 895A null detector, R202, C202, R203, and C203 form a double-section, low pass filter that reduces any AC component present on the DC voltage being measured. The difference between the voltage appearing at the output of the filter and the voltage across feedback resistor R208 is converted to a square wave by PC201 and PC202, an 84 Hz photo-chopper. The effect of a negative input voltage is to shift the phase of the square wave by 180°. Thus, by using a phase sensitive demodulator, the polarity of the input voltage is reconstructed in the meter drive circuit. The alternating voltage created by PC201 and PC202 is amplified by a five-stage amplifier. Transistors Q201, Q202, and Q203 are common-emitter amplifier shaving negative feedback. Transistor Q204 is a common-emitter amplifier which drives the push-pull pair Q205 and Q206. The output from the push-pull pair is applied to the primary of transformer T201. Transformer T201, which provides isolation between the null detector and the grounded recorder output, has two secondary windings , one of which is connected to null detector common, the other winding connected to chassis common. The output of the winding that is connected to chassis common is applied to a phase-sensitive demodulator, Q209 and Q210. The square wave drive for this demodulator is from a winding connected to chassis ground on the drive transformer, T101. The polarity of the square wave out of the transformer permits conduction of Q209 or Q210 during alternate halves of the square wave cycle. Since the phasing of the square wave  through the null detector depends on the polarity of the input voltage, then either Q209 or Q210 conducts, depending on the polarity of the input. Conduction of Q209 or Q210 tends to drive current through the meter in opposite directions. Thus, the meter and recorder output are driven according to the polarity as well as the magnitude of the unknown voltage.

3.11. The output of the secondary winding that is connected to null detector common is applied to another phase-sensitive demodulator, Q207 and Q208. The square-wave drive for these transistors is obtained from a winding that is connected to null detector common on T101. This phase-sensitive demodulator operates identically to Q209 and Q210. The output from Q207 and Q208 is filtered and fed back to R208 at the input of the null detector for gain stabilization.

3.12. The null detector has a basic sensitivity of 1 mV, except for the most sensitive null ranges on the 1 and 10 volt ranges, in which the sensitivity of the null detector is increased to 100 µV.

3.13 TVM attenuator. For null measurements the voltage difference of the unknown voltage minus the reference voltage is divided down as necessary by positions on the TVM attenuator as selected by null switch  sections S3E, S3G, S3K and S3I, which provides the basic null detector input of 1 millivolt.

3.14 Chopper Drive Circuit.  The chopper drive circuit provides AC drive for the photo-chopper modulator and for the phase-sensitive demodulators. Auxiliary voltages for the null detector are also obtained from the chopper drive circuit. The chopper drive is essentially a transformer-coupled multivibrator Q109 and Q110. Assume that Q110 has been conducting, and Q109 turns on. Conduction of Q109 applies approximately +18 volts between pins 1 and 2 of T101, which induces a voltage between pins 4 and 5. The secondary winding of pins 4 and 5 is connected between the bases of Q109 and Q110. The two windings are phased so that the voltage between pins 4 and 5, approximately 12 volts, adds to the voltage at the base of Q109, which is effectively clamped at +18 volts. The resulting voltage biases Q110 off. After Q110 is biased off, capacitor C112 begins to discharge through R145 and R146. When the voltage across C112 approaches zero volt, Q110 begins to turn on, and the preceeding cycle is repeated with Q110 conducting. Resistor R146 is used for frequency adjustment.

3.15. AC voltage for the +17 volt, +10 volt, and –17 volt auxiliary supplies is obtained from the winding of T101. Diodes CR206 and CR207, and capacitors C214 and 215, form two half-wave rectifiers. Regulation of the +10 volt supply is provided by zener diode CR203. Illumination for the photo-choppers PC201 and PC202 is also obtained from the same winding of T101, through CR204, CR205, DS201 and DS202.  Diodes CR204 and CR205 permits conduction of DS201 and DS202 on alternate half-cycles of the 84 Hz wave, thus illuminating PC201 and PC202 alternately. Since the resistance of the photocell decreases several times when illuminated, the result is the same as that of a mechanical chopper. Lamp DS202 also drives the photocell chopper PC101 in the reference supply.

3.16 Effect of AC components. The low pass filter at the input of the chopper-amplifier will attenuate any AC component present on the DC being measured. The magnitude of the AC voltage appearing at the output of the filter depends on its amplitude and frequency before filtering. The only AC voltage tat will reduce the accuracy of the 895A is one that either saturates the null detector or beats with the chopper frequency. Since the voltage required for saturation is greater than that required for beating, the null detector is most senstitive to an AC voltage having a frequency that is a submultiple or a low multiple of the chopper frequency. However, this is easy to detect, because the meter will oscillate at the difference frequency. For all practical purposes, no trouble should be encountered above a hundred Hz. If low-frequency AC components that affect accuracy ever occur, additional filtering in the operating instructions will eliminate the problem.